## Slide 1: Before we get started…
For your research,
  what is the frequency content of the signals you work with?
  what is their amplitude?
  how accurately do you need to measure the amplitude of a signal?
  how accurately do you need to measure the phase of a signal?

## Slide 2: Instrumentation and metadata: a brief overview
with material adapted from the Instrument Response Functions tutorial developed for USArray data processing course by Bob Woodward and Suzan van der Lee
Emily Wolin
USGS Albuquerque Seismo Lab

## Slide 3: Take-home points
Instrument responses are not a black box! They describe the process of taking a physical ground motion and converting it to a digital signal
Choose your processing parameters carefully – what do you need out of your data?
Understand the capabilities and limitations of the instrumentation you are working with
Seismic instrumentation literature can be dense and hard to read for non-specialists. I am always happy to answer your questions about metadata and instruments! ewolin@usgs.gov
We’re working on an SRL EduQuakes manuscript that discusses how to select the right instrumentation for your study: please let me know what you’d like to see!

## Slide 4: Why do *you* need to think about instrumentation and metadata?
Seismology’s history of open data sharing = data from a wide variety of seismic instrumentation is readily available to you as a researcher
Important to choose the right tool for the job
Important to understand instruments’ capabilities and limitations
Need to use metadata to convert raw signals to physical units

## Slide 5
Global Seismographic Network:
Highest-quality, long-term observations for research and monitoring
Special emphasis on long-period data (T>100 s)
Global coverage
ANSS backbone network:
Monitor M≥2.5 earthquakes inside United States
Provide uniform coverage across the lower 48, fill gaps between denser regional networks
Portable stations:
Rapid response to significant earthquakes
Densify coverage to improve estimates of earthquake location and depth
Record near-source strong ground motions for hazard studies
Notes: three examples of different kinds of stations with different instrumentation

## Slide 6: Outline
What is an instrument response?
Very brief overview of noise sources
Self-noise and instrumentation limitations
Follow up on questions from last week: channel naming
Exercise: removing instrument responses in ObsPy!

## Slide 7: From Earth to Your Desktop
seismometer
digitizer/ datalogger
volts
data acquisition
data management center
ground motion
miniSeed
miniSeed & metadata
waveforms
metadata
mechanical response
(poles and zeroes)
electrical response
(“gain factor”)*
AD response
(FIR filters)
counts
Notes: assumed linear – small non-linearities ignored
allowable for feedback seismometers.

## Slide 8
Stein & Wysession (2003)

## Slide 9
Stein & Wysession (2003)
Notes: Should really be sources (plural) because of course a seismometer records *any* earth shaking

## Slide 10: What is an instrument response?

## Slide 11
Output signal (seismogram in counts)
Input signal (ground motion)
=
Instrument response
so to remove the response and get a seismogram in units of ground motion…

## Slide 12: Removing instrument responses: Example from Unit 1
Deconvolving instrument response = division in frequency domain
Note that the amplitude response rolls off toward zero at low and high frequencies – deconvolution needs to be done in a way that avoids dividing by small numbers.
You will explore this later in the exercise!
(In reality, we are working with complex quantities, but it’s helpful to visualize using just the amplitude spectra.)
wait a minute, what’s going on here?

## Slide 13: From Earth to Your Desktop
seismometer
digitizer/ datalogger
volts
data acquisition
data management center
ground motion
miniSeed
miniSeed & metadata
waveforms
metadata
mechanical response
(poles and zeroes)
electrical response
(“gain factor”)*
AD response
(FIR filters)
counts
Notes: assumed linear – small non-linearities ignored
allowable for feedback seismometers.

## Slide 14
Three types of seismometers used within TA
Seismometer response specified by poles and zeros
For other sensors, check out the IRIS Nominal Response Library: http://ds.iris.edu/NRL/

## Slide 15: Obtain and plot a response object with ObsPy
import obspy
from obspy.clients.fdsn import Client
client = Client("IRIS")
t1 = obspy.UTCDateTime(‘2020-07-01’)
inv = client.get_stations(network="IW", station="PLID", channel="BHZ", level="response", starttime=t1)
inv += client.get_stations(network="GS", station="PR01", channel="HHZ", level="response", starttime=t1)
inv.plot_response(min_freq=1e-3)
inv.write(‘inventory.xml’, format=‘stationxml’)

## Slide 16: Metadata formats
StationXML, SEED (dataless, RESP), SAC PZ
  move toward StationXML: human-readable, more flexibility, more info about equipment (if operator supplies it)
Open inventory.xml in your preferred text editor and identify stages of the response:
  Seismometer (conversion from ground motion to volts, poles and zeros, sensor gain)
  Digitizer (conversion from volts to counts, digitizer gain)
  FIR filters (digitizer antialiasing – may be several rounds of these!)
Question from earlier chat: how do I tell if a response is specified in units of velocity, acceleration, etc?
  Response is usually specified in the sensor’s “native” units: broadband = velocity, accelerometer = acceleration… but not always. Check the response!

## Slide 17: The signal processing chain
High initial sample rate - no anti-alias filtering
Cascade of Finite Impulse Response (FIR) filters
  Provide anti-alias low-pass filters
Each low-pass followed by decimation
seismometer
digitization
analog signal (volts)
anti-alias FIR filter & decimate
anti-alias FIR filter & decimate
anti-alias FIR filter & decimate
5120 sps
(counts)
320 sps
80 sps
40 sps
BHZ @ 40 sps

## Slide 18: FIR Cascade
Combined effect of FIR cascade prevents aliasing
sample at 320 sps after filtering
sample at 80 sps after filtering
sample at 40 sps after filtering
sample at 20 sps after filtering

## Slide 19: FIR Cascade - Cumulative Response
Response shown for final BH channel sampled at 20 sps

## Slide 20: The TA: Composite FIR Filter
Represent the entire FIR filter cascade with a single filter
Seed provides time domain representation
Time Domain Representation
Frequency Response

## Slide 21: FIR Filter Ripple
Ripple near Nyquist Frequency
Ripple is less than 5%
Nyquist Frequency = half of sampling frequency

## Slide 22: Recap on FIR Filters
What are they used for?
  Low-pass anti-alias ﬁltering
When can you ignore them?
  When working at frequencies reasonably* below the Nyquist frequency
What are their side-eﬀects?
  Slight rippling (in amplitude) near Nyquist frequency
  Acausal ringing for sharp onsets (ringing is at frequencies near Nyquist)
What can be done about the side-eﬀects?
  Knowledge is power - steer clear of side-eﬀects
  My personal practice:
    max frequency analyzed should be no higher than 50-75% of Nyquist (i.e., 10-15 Hz for data sampled at 40 samples per second)
    just do the full deconvolution!
Notes: reasonably means:
if fNyq = 20 Hz (sps = 40 Hz)
10 Hz at 2+ % accuracy
2 Hz at 1- % accuracy
0.5 Hz at “fraction of a” % accuracy

## Slide 23: Outline
What is an instrument response?
Very brief overview of noise sources
Self-noise and instrumentation limitations
Follow up on questions from last week: channel naming
Exercise: removing instrument responses in ObsPy!

## Slide 24: Seismometers record signals from many sources
SPREE portable deployment (Wolin et al., 2015)
Transportable Array (Busby et al., 2018)

## Slide 25: Drop in “cultural” noise levels due to covid-19 lockdowns
Raspberry Shake installed in a school in Nepal
(@AtSeismo/Shiba Subedi)
Permanent broadband in Brussels
(@Seismologie_be/@seismotom/Thomas Lecoq)

## Slide 26: Seismometers record signals from many sources
(Ebeling, 2012)
(Koper and Burlacu, 2015)
Ocean storms – including hurricanes

## Slide 27: Long-period signals from surface tilt

## Slide 28: Long-period signals from surface tilt
Tytell et al. (2011)
Vernon et al. (2011)
Passage of storm front over TA stations:
large tilt signal coincides with passage of gust front (initial pressure jump)

## Slide 29: PSDs, PDFs, and noise models
Peterson (1993) noise models widely used to assess station performance
Defined by range of power spectra observed at high-quality permanent network sites
“New” High and Low Noise models (NHNM, NLNM)

## Slide 30: PSDs, PDFs, and noise models
Power spectral density (PSD) probability density functions (PDFs) show distribution of noise in a given time period
Slice seismogram into many overlapping segments, calculate power spectra, create 2D histogram of power vs. frequency
Widely used in network quality control
Choice of processing parameters matters! Tune carefully to avoid biasing or smoothing through signals of interest (Anthony et al., 2018)
(McNamara and Buland, 2004)

## Slide 31
short
0.1-1 s
1-10 Hz
microseismic
1-20 s
0.05-1 Hz
long
> 20 s
< 0.05 Hz
Human activity
Ocean storms
Local site response to atmospheric & thermal variations
PSDs, PDFs, and noise models
At T > 20 s, horizontal components are usually noisier than verticals due to local tilt effects
Horizontal noise decreases with increasing depth of installation
Site selection
Construction
Noise mitigated by:
Not much!

## Slide 32: Outline
What is an instrument response?
Very brief overview of noise sources
Self-noise and instrumentation limitations
Follow up on questions from last week: channel naming
Exercise: removing instrument responses in ObsPy!

## Slide 33: Self-noise
Defined by instrument + digitizer
Smallest resolvable signal based on limits of instrumentation
Instruments can have the same response shape but different self-noise levels!
  example: Trillium 120 vs Trillium Compact 120
Trillium 120
Trillium Compact
smaller masses in TC120 = higher self-noise at long periods!
Hutt and Ringler (2010)

## Slide 34
Self-noise
Example: the lowest noise levels recorded by the TA are defined by the instrument + digitizer self-noise floor, NOT physical ground motion
Tradeoff between self-noise and dynamic range: can use higher-gain sensors and/or 26-bit digitizers to detect smaller signals, but this lowers the clip threshold
(Wolin and McNamara, 2019)

## Slide 35
Self-noise
If you’re used to working with data from permanent broadband sensors, you may not think much about self-noise
Short-period systems (geophones, nodes) and accelerometers – especially MEMS – have self-noise that is much higher than a typical observatory-grade broadband sensor + digitizer
Know the limitations of your instruments!

## Slide 36: Outline
What is an instrument response?
Very brief overview of noise sources
Self-noise and instrumentation limitations
Follow up on questions from last week: channel naming
Exercise: removing instrument responses in ObsPy!

## Slide 37: Channel naming conventions
Orientation: why use Z, 1, 2 instead of Z, N, E components?
  Some stations are deployed in situations where they cannot be oriented to true North
    Posthole or borehole
    Ocean-bottom seismometer
  Have you ever installed a station and later realized its orientation was incorrect?
  Best practice: use orientation specified in metadata instead of assuming orientation based on channel name
  Determining TA orientations:
    Ekström, G., and M. Nettles, 2018, Observations of seismometer calibration and orientation at USArray stations, 2006–2015, Bulletin of the Seismological Society of America, doi: https://doi.org/10.1785/0120170380
    Ekström, G., and R.W. Busby, 2008, Measurements of seismometer orientation at USArray Transportable Array and Backbone stations, Seismological Research Letters, doi: https://doi.org/10.1785/gssrl.79.4.554

## Slide 38: Channel naming conventions
Interpreting channel codes:
1st letter: sample rate
H = 80-250 samples/sec
B = 10-80 sps
L = 1 sps
V = 1 sample/10 sec
2nd letter: instrument type
H = high-gain seismometer
(resp proportional to velocity)
N = accelerometer
3rd letter: orientation
N = north
E = east
Z = vertical
1, 2 = arbitrary horizontals
Z
H
L
Sensor proportional to velocity (usually broadband),
Vertical component,
sampled at 1 Hz
https://ds.iris.edu/ds/nodes/dmc/tools/data_channels/

## Slide 39: Channel naming conventions
1st letter: sample rate
H = 100 samples/sec
B = 40 sps
L = 1 sps
2nd letter: instrument type
H = high-gain seismometer
(resp proportional to velocity)
N = accelerometer
1
N
H
Accelerometer,
horizontal component,
sampled at 100 Hz
https://ds.iris.edu/ds/nodes/dmc/tools/data_channels/
3rd letter: orientation
N = north
E = east
Z = vertical
1, 2 = arbitrary horizontals
Interpreting channel codes:

## Slide 40: Channel naming conventions
1st letter: sample rate
H = 100 samples/sec
B = 40 sps
L = 1 sps
2nd letter: instrument type
H = high-gain seismometer
(resp proportional to velocity)
N = accelerometer
E
N
H
Accelerometer,
East component,
sampled at 100 Hz
https://ds.iris.edu/ds/nodes/dmc/tools/data_channels/
3rd letter: orientation
N = north
E = east
Z = vertical
1, 2 = arbitrary horizontals
When in doubt,
INSPECT THE RESPONSE
Interpreting channel codes:

## Slide 41: Exercise: Obtain and remove instrument response
Processing parameters should be chosen based on instrumentation used and frequency band of interest for your study.
Open-ended lab: experiment with pre_filt, water_level, taper, and zero_mean parameters
  How do different pre_filt and water_level parameters affect the output?
  What happens if you add an offset or trend to your data and don’t taper or remove the mean?
Obtain data for your station and signals of interest
If you want more: try computing PSD PDFs for a station of interest to you following the ObsPy tutorial: https://docs.obspy.org/tutorial/code_snippets/probabilistic_power_spectral_density.html